How to calibrate a NIR2.5

Materials to use:

2x FC-IR600 fiber

1x Avalight-HAL-S

1x Avalight-CAL-AR

1x CUV-UV/VIS

1x STARNA Liquid type RM-NIR

Calibrating a NIR2.5 step by step:

Connect the HAL-S->IR600->CUV-UV/VIS->IR600->NIR2.5

Connect the NIR to the computer and start up the HAL-S.

Open Avasoft.

Start the measurement.

Push the SAC button to get the maximum output of the HAL-S trough the two IR600 fibers and the CUV-UV/VIS with the cover on top.

Close the shutter of the HAL-S and take a correct for dark.

Open the shutter and take a white reference.

Put the STARNA RM-NIR in the CUV-UV/VIS and close the cover again.

Now go to the ABSORBTION option for the first part of the calibration.

Now we see some peaks that we can use in the calibration.

Load the lit file: TS5_NIR_customized

Peaks you could use are:

When you have selected these peaks press calculate and it should be identical.

To finish the calibration you need to go to scope mode and use the ARGON lightsource.

Only 1 ir600 fiber is connected directly between lightsource and NIR.

Calibrate with argon lit file.

After the calibration is done safe it to Eeprom and make a new calibration report and stickers.